## Slide 1
Last Results from the 6T Fringe Tracker Test Bench
Configuration:
- Chip VLC-04-01 (2nd Run) – 6T ABCD
- 10 spectral channels between 1.5mm to 1.7mm
- Frame Time 8ms
- Actuators: Fast Piezo, stroke 33um
- Algorithm of GRAVITY (without Kalman)
- Servo Loop: pure integrator - Gain: =0.05 & =0.1
Results:
- Instrumental Phase Closure +/-4mm (above the specification +/-1mm)
        See plot GD Closure
- We can close the full loop (GD+PD)
        See plots PD-RefPD GD-RefPD
Next Steps:
- Inject perturbations in the system (directly on the fast piezo)
- Determine the good gains
- Test the code with MIRCx in lab (STS)
- Test the code on Sky with MIRCx and SPICA-FT

## Slide 2
=0.05 & =0
==0
=0.05 & =0.1