## Slide 1
SPICA-VIS Design Review, July 23-24, 2020
Optical study of the spectrograph
Cyril Pannetier
Requirement | Values
Number of telescopes | 6, linear non redundant configuration
Spectral range | 600 nm – 900nm (guaranty 650 – 850nm)
Spectral resolution | 140, 3000, 10000
Photometry needed | Photometric channels are needed to correctly process the visibilities.
Fringe measurement | We want to measure the complex visibilities of the 15 interferometric patterns.
Notes: Objective: measure the fringes of the 15 baselines due to the 6 telescopes. For that, we align the six fibers in direction and we disperse it in the other direction. We’ll do it with three spectral resolutions: 140, 3000and 10000. And we need a calibration for both telescope flux, so we need a photometric channel on the same detector than the interferometry. We observe a spectral range which goes from 600nm to 900nm. Due to the classical limit of throughput close to 650nm for optical elements, we want to guaranty 650nm (H-alpha ray) until 850nm (calcium triplet).

## Slide 2
SPICA-VIS Design Review, July 23-24, 2020
Optical study of the spectrograph
Cyril Pannetier
Detector
V-groove output
FSL
SSL
TSL collimating lens
SCL
FCL
Imaging lens CAO
Image of the µlenses
Fringes in focal plane
Beam splitter 10/90
V-groove
Detector
Pupil plane
50mm
Fibers outputs
Notes: Both arm is approximatly 1meter-long.
The light is first split in to parts by a 10/90 beam splitter and then both channel see different optical elements until reaching the dispersing part, which means from the diffracting grating to the detector.
At the top, we have the photometric channel: with the achromatic lens FLS and the SLS whose role is to reimage the fibers output behind L2 with the sizes that we want on the detector. Then the achromatic lens L3, with the achromatic lens L4 play the role of the collimator of the spectrograph.
At the bottom, the interferometric channel sees only two optical elements. The first FCL cylindrical lens creates an image that the Second Cylindrical Lens collimates before the spectrograph dispersion. This pair of cylindrical lenses performs the anamorphosis necessary for reaching a good sampling on the detector and the second cylindrical lens plays the role of the collimator of the spectrograph.
Here I show you the high resolution mode so the spectrograph is only composed of a diffraction grating and an imaging lens CAO.
For the low resolution mode, the grating is replaced with a flat mirror and a dispersin double prism makes the dispersion.
We are considering adding a tip-tilt mirror in order to make the alignment easier.

## Slide 3
SPICA-VIS Design Review, July 23-24, 2020
Optical study of the spectrograph
Cyril Pannetier
Fibers output – V-groove configuration
B=500µm
Configuration: 3B – 2B – 4B – 7B – B
Frequencies (B units):
1 – 2 – 3 – 4 – 5 – 6 – 7 – 8 – 9 – 11 – 12 – 13 – 14 – 16 – 17
To avoid crosstalk between
and :


Fiber coupling:
B
D
No crosstalk
Crosstalk
Notes: The camera we will use is a EMCCD (electron-multiplied couple-charge detector) with 1024x1024 pixels of 13µm size. With the design I presented, we reach sampling on the detector with 100 pixels per 1024 for the photometry, around 400 x 1024 for the interferometry. This enables us to keep a good readout rate (20Hz) which means 50 images/s, tipically the minimum readout rate to keep good contrast when fringe tracking is not working. At the bottom right we see the gaussian field propagation performed with Zemax.
Concerning the fringe sampling, the shortest one is around 2,94 pixels. The longest one on 65 pixels. The sampling in the sepctral direction and both direction for the photometric channel is around 2 pixels for the FWHM (full-width at half maximum).

## Slide 4
SPICA-VIS Design Review, July 23-24, 2020
Optical study of the spectrograph
Cyril Pannetier
Sampling on the detector
Camera: iXon Andor 588
Image sizes:
Photo:
Interfero:
  Total:
   50 images/s (20Hz)
Sampling:
  Fringe sampling:
    Shortest fringe on 2,94 pixels (@650nm)
    Longest fringe on 65 pixels
  Other sampling:
384pix
100pix
Interfero
Photo
Photo
Interfero
Spot Diagram des 12 voies
Interferometry:
Photometry:
Notes: The camera we will use is a EMCCD (electron-multiplied couple-charge detector) with 1024x1024 pixels of 13µm size. With the design I presented, we reach sampling on the detector with 100 pixels per 1024 for the photometry, around 400 x 1024 for the interferometry. This enables us to keep a good readout rate (20Hz) which means 50 images/s, tipically the minimum readout rate to keep good contrast when fringe tracking is not working. At the bottom right we see the gaussian field propagation performed with Zemax.
Concerning the fringe sampling, the shortest one is around 2,94 pixels. The longest one on 65 pixels. The sampling in the sepctral direction and both direction for the photometric channel is around 2 pixels for the FWHM (full-width at half maximum).

## Slide 5
SPICA-VIS Design Review, July 23-24, 2020
Optical study of the spectrograph
Cyril Pannetier
Interferometric channel
310 pixels @95%E
Image patterns of the interferometric channel in the spatial (left) and spectral (right) directions.
2 pixels @FWHM
Notes: Here is what the gaussian propagation simulation on Zemax predicts for the interferometric channel. We expect that 95% of the energy is spread over 310 pixels.

## Slide 6
SPICA-VIS Design Review, July 23-24, 2020
Optical study of the spectrograph
Cyril Pannetier
Photometric channel
Expected position of the closest photometric beam
4 pixels
4 pixels
Interferometric pollution:
< 0,5% of the photometric energy

## Slide 7
SPICA-VIS Design Review, July 23-24, 2020
Optical study of the spectrograph
Cyril Pannetier
Mechanical implantation
~320mm
capteur
Lentille imageante
R10000
R140
R3000
Heavy proposition
50mm
Prism on the wheel
 Automatic selection of the low spectral resolution

## Slide 8
SPICA-VIS Design Review, July 23-24, 2020
Optical study of the spectrograph
Cyril Pannetier
Mechanical implantation
~320mm
capteur
Lentille imageante
R10000
R140
R3000
Heavy proposition
50mm
Prism on the wheel
 Automatic selection of the low spectral resolution

## Slide 9
SPICA-VIS Design Review, July 23-24, 2020
Optical study of the spectrograph
Cyril Pannetier
Mechanical implantation
~220mm
capteur
Lentille imageante
R10000
R140
R3000
50mm
Light proposition
Prism outside the wheel
 Manual selection of the low spectral resolution

## Slide 10
SPICA-VIS Design Review, July 23-24, 2020
Optical study of the spectrograph
Cyril Pannetier
Mechanical implantation
~280mm
capteur
Lentille imageante
50mm
R10000
R140
R3000
Medium proposition
Prism on the wheel
 Automatic selection of the low spectral resolution

## Slide 11
SPICA-VIS Design Review, July 23-24, 2020
Optical study of the spectrograph
Cyril Pannetier
Low spectral resolution: R=140
Spectral dispersion:
Two commercial equilateral dispersing prisms made of F2: D=50mm

## Slide 12
SPICA-VIS Design Review, July 23-24, 2020
Optical study of the spectrograph
Cyril Pannetier
Medium and High spectral resolution
We need two diffraction gratings of 50x50mm.
For R=3000:
For R=10000:
Newport: blazed for 500nm and 550nm.
Efficiency
We expect better with Jobin-Yvon, our contact is in holidays.

## Slide 13
SPICA-VIS Design Review, July 23-24, 2020
Optical study of the spectrograph
Cyril Pannetier
Fiber Back Illumination
Collimated anamorphosed laser source
Movable mirror
Back-propagating beam shape
Back-propagating beams
Toward photometric chanel
Toward interferometric chanel
Figure 23: Optical concept of the back fiber injection system
100mm

## Slide 14
SPICA-VIS Design Review, July 23-24, 2020
Optical study of the spectrograph
Cyril Pannetier
Low resolution:
  [650-850] on 100 pixels
Medium resolution:
  Bandwidth: 70nm @650nm
High resolution:
  Bandwidth: @650nm
Throughput: > 86%  95% ?
Max Readout rate: 20Hz (50 images/seconde)
Performances

## Slide 15
SPICA-VIS Design Review, July 23-24, 2020
Optical study of the spectrograph
Cyril Pannetier
Dimensioning of the system
Element | Focale (mm) | Espace avant (mm) | Diamètre (mm) | Détails
Microlentilles (NA=0,12 = NA_fibre)  | 0,900 | 0,900 | 0,218 | Custom AMUS.de
Beam-Splitter |  | 200 | 12 | Transmission 10/90 Custom
Element | Focale (mm) | Espace avant (mm) | Diamètre (mm) | Détails
 | 0,900 | 0,900 | 0,218 | Custom AMUS.de
Beam-Splitter |  | 200 | 12 | Transmission 10/90 Custom

## Slide 16
SPICA-VIS Design Review, July 23-24, 2020
Optical study of the spectrograph
Cyril Pannetier
Dimensioning of the system
Element | Focale (mm) | Espace avant (mm) | Diamètre (mm) | Custom ?
Miroir M1 |  | 35 | 12 | Coating
L1 | 100 | 65 | 25,4 | x
L2 | 15 | 115 | 6,25 | x
L3 | 500 | >500 | 25,4 | x

## Slide 17
SPICA-VIS Design Review, July 23-24, 2020
Optical study of the spectrograph
Cyril Pannetier
Dimensioning of the system
Element | Focale (mm) | Espace avant (mm) | Diamètre (mm) | Custom ?
L1 cyl | 9,69 | 30 | L=12; H=10 | Non
L2 cyl | 400 | >400 | L=20; H=40 | Oui

## Slide 18
SPICA-VIS Design Review, July 23-24, 2020
Optical study of the spectrograph
Cyril Pannetier
Dimensioning of the system
Element | Focale (mm) | Espace avant (mm) | Diamètre (mm) | Custom ?
Miroir M2 (ou réseau) | - | 50 | 50,8 | Coating
Prisme | - | 100 | 50,8 | Oui
Lentille imageante L4 | 500 | 50 | 50,8 | Oui